Synopsis of Chapter 20: Thermal Control from

Building Electro-Optical Systems, Making It All Work

By Philip C. D. Hobbs

Doug Essex

62 N. Worth Ave.

Elgin, IL 60123

dessex@email.arizona.edu

# Introduction

This chapter is a freely available pdf file from the authors website (www.pergamos.net) of a chapter that was cut from his book “Building Electro-Optical Systems: Making It All Work” (Wiley, 2000).  The stated purpose of the book is to collect the ‘lore’ i.e. the hard earned experience of what is actually important when building electro-optical systems.  The book has a delightful mix of informal writing and well thought out explanations with enough of the derivations of mathematical relationships to know the limits of applicability and point you toward further depth if desired.  Interleaved throughout are asides about how real components and systems behave, as opposed to the idealized versions more familiar from introductory classes.  The book looks at optics and realistic optical components in detail, and at detectors and their associated electronics and signal processing, but tends to gloss over optomechanical topics, with the exception of this chapter.

Hobbs emphasis in this chapter is thermal control, i.e. actively heating or cooling components to keep them at their nominal operating point versus what has been the emphasis in OPTI 521, thermal compensation or the passive use of materials to minimize thermal effects.  Of course in any real system both approaches need to be used.  Thermal compensation can ease the requirements for the thermal control system, making it simpler and more reliable.

Hobbs gives some graphic examples of the effects of thermal expansion and thermal gradients and points out that gradients are usually more difficult to deal with.  One piece of advice paraphrased from E. Loewen, a well known worker in the field of diffraction grating ruling engines (famous for the precision required for their motion) is to always look for bending of beams due to thermal gradients first when tracking down thermal problems.

After a basic development of thermal transfer mechanisms (conduction in solids, radiative, conduction thru gases and convection he looks at methods for achieving uniform air temperature – use a box and stir the air with a fan.  He references measurements showing that a simple cardboard box and fan had nonuniformities <3 mK at room temperature.  He recommends the use of Styrofoam or vacuum dewars for insulation, giving a literal rule of thumb for selection of superior grade Styrofoam “Use the soft stuff… that squeaks when you rub hard with your finger, the harder stuff … that crunches instead of squeaking is inferior.” He has a rather long aside about the dangers and remedies for condensation on cooled surfaces – dry gases from various sources are a good starting point, but seals must be hermetic or atmospheric moisture will penetrate.  Desiccants can be used but are much better cooled than simply dumped in.  They also tend to be particle generators that can be as troublesome as condensation.  He mentions an interesting trick for getting dry air for one of and prototyping work.  Since cold air carries far less water than warm, place the pieces in a chest type freezer until they equilibrate, then open and assemble.  The chest will prevent the dry air from escaping when opened.  I would like a Plexiglas glove panel too.  As an aside from my own experience, it is almost always better to use really dry air than an inert gas.  Over time organic compounds, even “space qualified” ones can outgas.  In high power systems this organic contamination can be photophoretically concentrated on optical surfaces where the fluence is highest, resulting in catastrophic damage.  The oxygen in the air will burn off these organics and prevent buildup.  This trick has solved problems in systems as diverse as military aerospace to submarine repeater links.

Hobbs then gets into the meat of the chapter – how to build thermal control systems.  At least a nodding familiarity with control system ideas like loop gain, proportional control, is beneficial to getting the most out of this chapter.  In order to control anything, you first have to have a measurement and measurements require sensors.  Hobbs covers the good and bad points of most precision temperature measurement devices – things like thermometers and thermostats are given short shrift since thermometer data is seldom electronically read out and used elsewhere and thermostats if electronically read out are already part of on-off (bang-bang) controllers.  His favorite sensors are linearized glass bead thermisters and Pt RTD’s.  They are stable, linear, fast response (100’s of ms).  His next favorite are IC sensors – also stable, linear devices, but much slower, limited to about 100°C on the high temperature side and they can dissipate enough heat to influence their surroundings.  Diodes are recommended as cheap, low accuracy sensors.  He gives a way to use the photodiode monitor output of a laser diode to also measure temperature – it’s always nice to get 2 measurements for the price of one.  Surprisingly, given the prevalence of thermocouples, we are told to “avoid them like fleas”.  Their nonlinearity is harder to correct – but they dissipate no power, are small and can respond very quickly.  Given that there many linearized thermocouple measurement systems available (my personal DMM has a plug and readout for a TC) I’m not entirely sure he is justified in this.  In closing he does recommend having a thermal switch to protect against overheating if the system dissipates more than 50W.  Later in the chapter he asks the questions about how well the sensor reflects the quantity that you actually want to control

Since what we always have is the temperature of the sensor, which can have some offset and lag from the temperature we actually want to control.

Once we’ve measured something we can add or take away to create the conditions we want – for thermal control this means heaters and coolers.  Providing heat is generally pretty simple – dissipate a little electricity, either through a resistor or the hot side of a thermoelectric cooler (TEC) – making sure this heat is transferred efficiently and uniformly is generally more of a problem.  Cooling can be as simple as using the other side of the TEC or as complicated as cryogenics, but he is excluding most cryogenics from consideration.  He does mention the use of mechanical refrigerators, but only briefly mentions the significant costs of refrigeration in size weight and power dissipation.  He goes into the limitations of TEC’s in some detail.  Since the hot and cold sides of a TEC are only a few millimeters apart, so convection from the hot side, conduction through the thermocouple elements and power dissipation all conspire to impair the efficiency of the TEC.  They have a fairly large thermal expansion between the hot and cold plates so they should not be rigidly mounted.  Repeated extreme temperature cycling can damage less expensive TEC’s.  Metal mounting screws can also be a thermal short, leaking heat into the cold plate, so mounting screw size should be minimized, or nylon screws used.  Thermal conduction and power dissipation of signal and power leads to the device on the TEC need to be considered also, very fine copper or gold wire or flexible circuit board (copper on Kapton) thermally contacted to the plate is generally recommended.  To operate at large DT’s we need a good heat sink with active cooling and excellent thermal contact.  He recommends lapping the sink with fine valve grinding compound and an old TEC.  Some useful rules of thumb are given for heat sink selection and mounting later in the chapter.  Natural convection heat sinks have a thermal resistance inversely proportional to their volume – but installation details matter a great deal, the fins should be oriented vertically and a box can really be a problem.  Forced air cooling can improve thermal resistance by a factor of 5-10.  Heat sink paste is actually not a great conductor so it should be kept as thin as possible.  Solder or silver epoxies are much better, but the epoxies must be cured per their manufacturers instructions or you probably won’t get the performance desired.  Thermal spreader plates of aluminum or copper are useful to avoid gradients due to the nonuniformity of most heaters and coolers.  Multistage TEC’s without spreaders have destroyed themselves with a too quick turn on, allowing large gradients to build up and melt connections.

One interesting method of increasing the capacity of the TEC is to use phase change cells (small amounts of material that undergo a phase change close to the operating temperature).  These cells will absorb a great deal of energy when the material inside melts.  This is useful if the device on the cold plate has a low duty cycle and high peak heat loads.  The phase change cell handles the peak and the TEC keeps things at a steady state

A simplified model for a TEC is given to enable effective controllers to be designed is given.  The heat flowing into of the cold plate is given by

and heat flowing into the hot plate is given by

All of these are parameters can be obtained from the TEC datasheet.

When the control system has to deal with large transients we need as large a bandwidth as we can get.  Two crucial bandwidth limiters are slow diffusion and large thermal mass.  To combat slow diffusion the sensors are generally placed in close proximity to the heater/coolers, spreader plates help keep the temperature the sensor is measuring close to the temperature of the portion of the instrument we actually care about.  Common centroid designs – layouts with symmetrical positions of heaters/coolers and sensors around the part whose temperature we are trying to control is a good way to cancel gradients.

# Conclusion

This chapter is very representative of the book as a whole – filled with useful practical tricks, but never oversimplifies the underlying physics.  It’s also some of the most fun technical reading you’ll ever do.  Attached below is the pdf file of the chapter.

# Endnotes

Hobbs, P.C.D. ,Building Electro-Optical Systems: Making It All Work , Wiley (2000)

pdf of this chapter is available at www.pergamos.net

Sharps, J. A. “Packaging Induced Failure of Semiconductor Lasers and Optical Telecommunications Components”, 27th Annual Boulder Damage Symposium: Laser Induced Damage in Optical Materials: 1995, SPIE Volume 2714 (1996) p. 676

Jollay, R. A., “Manufacturing Experience in Reducing Environmentally Induced Failure of Laser Diodes”, 27th Annual Boulder Damage Symposium: Laser Induced Damage in Optical Materials: 1995, SPIE Volume 2714 (1996) p. 679