Requirements for Fluidic Lens to be used in Opthalmic Correction

Randall Marks

- System Performance Requirements

  - Lens must fit in front of the human eye

    - Should be located at the front focal point of the eye.

    - Horizontally, edge of the lens must be less than 1” from the center of curvature to fit in front of the eye.

  - Lens must be capable of producing Zernike terms Z[2,0], Z[2,-2], and Z[2,2] independently.

    - Lens must be capable of producing 3 diopters of astigmatism in any direction.

    - Lens must be capable of producing ±5 diopters of power.

    - Lens must be capable of repeatably and consistently producing changes of less than 0.25 diopter of astigmatism and power.

  - Lens must have a 6-32 x 1/8” screw hole for mounting in the test equipment.

  - Individual lens cells must be individually testable.

  - Combination of alignment and lens error must produce less than 0.30 microns RMS error with a 6mm diameter at maximum operating deflection on axis to be unnoticeable to the human eye within a safe margin.

  - Design should incorporate prior design technology to better predict performance and expedite fabrication

  - Lens stack should be no thicker than 2 cm, and, to improve field performance, should be as thin as possible.

- Lens Construction Requirements (from prior design preference)

  - Must have a 0.5mm hole in the structure for injecting fluid into the lens cell.

  - Must have a 0.5mm hole in the structure for bleeding off excess gas pockets.

  - PDMS membranes are to be used for lens surfaces, which require power.

    - The membrane front surface must be formed against a 2” diameter, 0.25 wave glass flat, and must be flat to gravity within 0.25°.

  - The lenses are to be filled with deionized water.

  - Each membrane must be aligned with a minimal amount of tension in all planar directions from the center.